Ocean Optics Hyperspectral radiometer

# Above-water reflectance

## Introduction

## This instrument will measure water-leaving remote-sensing reflectance (RRS) which is determined as the water-leaving radiance (LW) over the downwelling irradiance (ED).

RRS =LW/ED

## However, the measured radiance (LM) by the above-water radiometer is compose of the water-leaving component and the light reflected on the surface of the water as well.  To eliminate the last component, Mobley (1999) has proposed to measure the downwelling light (LSKY), with respect of the zenith angle (30o) and to multiply this signal to a “reflectance factor ”.  The water-leaving radiance is then calculated as:

## LW = LM - LSKY

## The downwelling irradiance is measured using a Spectralon grey card, which diffuse perfectly the incident light.

## To make this remote-sensing reflectance we then need three types of measurements, a sea measurement, a spectralon measurement and a sky measurement.  As it is pointed out in the NASA protocol, the light measurements should be made at an azimuth angle of 90o to 180o from the sun to reduce the

## reflected light component in the measured light.

### Protocol

#### Set up of the instrument

- Insert the optic fibre (2) in the support (1) and plug the other end in the hole at the front of the Ocean Optics reader (3).

- Connect the Ocean Optics reader (3) to the lap top (7).  Connect the lap top to the power (8).

- Connect the Ocean Optics reader (3) to the power (4 and 5) (back side of the reader) following the order that is indicated on the wire (top first and bottom second).  If the Ocean Optics reader is functional, a red light will appear in the front side of the reader, it will turn to green when the temperature has reached ~20oC.  If the Ocean Optics reader is not connected properly, a flashing red light will show up, disconnect the reader and reconnect it to the power by inserting first the wire at the top of the reader and then the one at the bottom.

- Turn on the computer (7) and double-click on the Ocean Optics icon.  The system is now installed.

- You will have to acquire data from the sea.  Then you take data from the Spectralon card (9) to get the downwelling irradiance.  To get this, you install the Spectralon support (10) in the holes of the optic fibre support (1) and make sure the optic fiber is really pointing at the Spectralon card (9).  You will also take measurements of the light from the sky to get the light that is reflected on the surface of the ocean.  To get these measurements, you take off the Spectralon support and you turn the optic fiber support up-side down, make sure that you respect the angles, zenith angle of 30o from the nadir direction and the same azimuth angle than when you made the other measurements (between 90o and 180o away from the sun).  To get these data, you will have to put a 1% filter in front of the optic fibre to reduce the incident light.

4

6

5

2

1

8

3

7

10

9

## Acquisition of the data

- Once the program is turned on, configure the acquisition.  Click on “Time Acquisition”, then “Configure” and then “Configure Acquisition”.  A “time Acquisition configuration” window will appear.  You have to change the name of the files (month, day, year).  If you want to change the place to save the file (ex: leg2 instead than leg1) you press on the “…” button next to the filename window and change the pathway and add the filename and press save.  Make sure that you “save full spectrum with each acquisition” and that you “save every acquisition”.  These options should already be selected.  Press “OK”

- Activate the time acquisition. Press “Time Acquisition” and then “Activate Time Acquisition” or press the clock located directly over the “integration time” window in the toolbar.  The system is now ready to collect the data.

- Change the magnitude of the spectrum.  Before you collect the data, make sure that you have a spectrum that use the maximum of the available magnitude without capping at 4000.  The maximum of the spectrum should reach about 3000-3500.  To change the magnitude of the spectrum, change the “integration time” which is located directly over the graph. The integration time on a sunny day is usually about 40-50 msec for the sea measurements, about 15-25 msec for the Spectralon measurements and about 150 msec for the sky measurements (after the 1% filter is add in front of the optic fibre). The minimum is 3 msec

- Make sure that the “average” is 5.  Each spectrum will be the average over five spectra.

- Make sure the spectrum is “Corrected for the Electrical Dark”.

- You are ready to acquire the data.  Click on “Time Acquisition” and then “Start” or click on the green triangle in the toolbar (over the integration time).  You will see the number of the acquisition files at the bottom of the spectrum.  Take about 7-10 files for each type of measurements (sea, Spectralon, sky).

- Change the type of measurements.  Click on “Time Acquisition” then “Pause” or on the “pause button” of the toolbar (next to the start button) and change your set up (add the Spectralon or turn the support up-side down and change the “integration time”) then press “Pause” again, the files will increment starting where it has previously stopped. Don’t click on “Stop” until you are finished because if you start again without changing the configuration of the files, it will over-write on the files you have previously acquire.

- Note the number of files that you have acquire for each type of measurement and the time that you have started and finished the measurements

If you have any questions: Marie-Helene Forget

(902) 426-4681

forgetmh@mar.dfo-mpo.gc.ca